## Slide 1: Dual e31: Direct and Converse e31 Measurement for a Single Thin Film Sample
Sean W. Smith, Bernardo Martinez-Tovar, Michael S. McDaniel, Naomi B. Montross, Eduardo Villareal, and Joe T. Evans
Radiant Technologies, Inc.
2024 IEEE UFFC JS
2024 09 24
Booth 41

## Slide 2
Issue
  The effective transverse piezoelectric coefficient (e31) is useful for piezoMEMS design.
  Measuring e31 from the converse piezoelectric effect is straightforward.
  Measuring e31 from the direct piezoelectric effect has been more difficult.
  Radiant proposes our Dual e31 technique to measure both the converse and direct e31 simultaneously from a single sample
Tester
DRIVE
RETURN
Stimulus
Converse
Uniform strain response
Tester
DRIVE
RETURN
Response
Direct
Stimulus: strain
may be nonuniform
Notes: Capacitor generates uniform strain along the cantilever
Mechanical deformation doesn’t necessarily create a uniform strain. Modeling this can be challenging.
May require a different sample than for the converse measurement

## Slide 3
Two Capacitor Cantilever
Measure Point
Clamped
Top Electrode → Red Bottom Electrode → Blue
  75 mm long, 5 mm wide
  One capacitor to actuate → Converse e31
  One capacitor to sense → Direct e31
  Capacitor design:
    Mechanically coupled – single cantilever
    Electrically isolated – separate bottom and top electrodes
    Extend past measure point – uniform strain – cantilever forms arc
    Equal length and area – either capacitor can sense/actuate
    Capacitor thickness 1 µm << cantilever thickness 525 µm – use silicon mechanical properties
    Symmetric about long axis – minimizes twisting
Notes:   This is the layout of our two capacitor Dual e31 cantilever

## Slide 4
Dual e31 Measurement
z
x
tSi
tCap
  Apply voltage to the actuator capacitor to bend the cantilever.
  Measure the displacement.
    Converse e31 and surface strain
  Capture the charge generated by the sense capacitor.
    Sense polarization
  Calculate the Direct e31 coefficient by dividing the surface strain into the sense polarization.

## Slide 5
PZT Thin Film Dual e31 Cantilevers
  Silicon substrate (100) 525 µm thick with 500 nm thermal SiO2 and 40 nm TiO2
  Pt bottom electrode 150 nm lift-off
  ~1 µm PZT by sol-gel
    20/80 4% Nb
    30/70
    40/60
    52/48
    65/35 9%La
  Pt top electrode 100 nm lift-off
Cool Cantilever photo
Dual e31 Cantilever

## Slide 6
Test Fixture
  Built off our converse e31 measurement system
  Photonic displacement sensor self-calibrates
  Five BNC connections
    Four pogo pins on cantilever frontside
    Fifth for cantilever backside to ground substrate
    Vision controlled connection swapping in final design (Multiplexer)
  Lid (not shown) reduces thermal and electrical noise
Prototype dual e31

## Slide 7
Measure Each Capacitor
Tester
DRIVE
RETURN
  Drive outer capacitor
  Measure displacement
  Measure outer capacitor polarization
  Drive inner capacitor
  Measure displacement
  Measure inner capacitor polarization
Tester
DRIVE
RETURN

## Slide 8
Tester
DRIVE
RETURN
GROUND
GROUND
Measure Dual e31
  Drive outer capacitor
  Measure displacement
  Measure inner capacitor polarization
  Slow, 0.5 Hz (2s) or less
  Ground opposing electrodes
  Converse e31
  -14.4 C/m2
  Direct e31
  -11.9 C/m2
  Shapes match
Notes: NOTE: The Direct e31 response is inverted to the displacement loop!

## Slide 9
  Invert sense polarization and plot over displacement
  Sense polarization is linearly proportional to the surface strain, which is linearly proportional to stress, which is linearly proportional to the displacement
  Quantify by plotting sense polarization versus displacement.
Normalize and Overlay Results
Normalized Direct e31 Charge
Normalized Displacement Loop
Sense Polarization:
Displacement:

## Slide 10
  Good linear fit
  R2, coefficient of determination, lets us quantify fit.
  Suggests sense polarization due to displacement
  Poor fit could indicate other factors generating charge
    Mechanical vibration
    Thermal variation
    Leakage between capacitors
Plot PSense vs Displacement z
Sense Polarization (µC/cm2) versus Displacement (µm) for a 52/48 PZT coupled e31 measurement.

## Slide 11
Results vs Composition
PZT | e31 (C/m2) |  | R2 Fit
 | Converse | Direct | 
4%Nb 20/80 | -12.3 | -7.7 | 0.9992
30/70 | -13.4 | -10.5 | 0.9989
40/60 | -9.8 | -9.5 | 0.9995
52/48 | -14.4 | -11.9 | 0.9975
65/35 | -9.8 | -3.8 | 0.9970
9%La 65/35 | -8.2 | -1.4 | 0.8829
  Direct e31 < Converse e31
  Most compositions have a good fit
  53/47 e31 in literature [1]
    Converse -15.0 C/m2 (at 10 V/µm)
    Direct -6.4 C/m2
[1] Yuichi Tsujiura, et al., Jpn. J. Appl. Phys., 54, 10NA04 (2015).

## Slide 12
Results vs Composition
PZT | e31 (C/m2) |  | R2 Fit
 | Converse | Direct | 
4%Nb 20/80 | -12.3 | -7.7 | 0.9992
30/70 | -13.4 | -10.5 | 0.9989
40/60 | -9.8 | -9.5 | 0.9995
52/48 | -14.4 | -11.9 | 0.9975
65/35 | -9.8 | -3.8 | 0.9970
9%La 65/35 | -8.2 | -1.4 | 0.8829
Chart: Bias (V)
  Converse follows displacement trend
  Direct follows sense polarization trend
  Fit poor for very small sense polarization

## Slide 13
  Dual e31 measurement simultaneously provides converse and direct e31 coefficient from a single sample.
  Cantilever design simplifies applied strain measurement.
  Linear relationship between displacement and sense polarization as expected.
  The Dual e31 fixture will be available fall 2024.
  Come visit us at booth 41 for a demonstration.
Summary
Sense Polarization (µC/cm2) and Displacement (µm) vs. Applied Bias for a 52/48 PZT coupled e31 measurement.
Notes:   Dual e31 measurement simultaneously provides converse and direct e31 coefficient from a single sample. And we demonstrated this for a variety of compositions of PZT
  Cantilever design simplifies applied strain measurement. Having two electrically independent capacitors is key
  Linear relationship between displacement and sense polarization as expected. Deviations from this could be a sign of noise or other outside effects on the measurement.

## Slide 14: Extra slides

## Slide 15
Test Procedure
  Characterize and orient the actuator capacitor.
  Characterize and orient the sense capacitor.
  Drive the actuator capacitor against ground while capturing the cantilever displacement and simultaneously measuring the charge generated by the sense capacitor referenced to ground.
  Invert the sense capacitor loop then overlay the displacement and charge loops after normalizing them to unity.
    Do they overlay? If yes, the measurement is valid.
  Calculate the Converse and Direct e31 coefficients.

## Slide 16
Assumptions and Constraints
  Capacitors extend past measure point
    Constant surface strain → cantilever forms an arc (circle segment)
  Capacitor thickness (tCap) << Cantilever thickness (tSi)
    1 µm << 525 µm
    Use silicon mechanical properties
  Then with cantilever length (x), vertical displacement (z) and cantilever thickness (tSi) we can calculate the surface strain
z
x
tSi
tCap

## Slide 17
  Apply voltage to the actuator capacitor to bend the cantilever while measuring the displacement.
    Converse e31 and surface strain
  Simultaneously capture the charge generated by the sense capacitor as it is bent by the actuator capacitor.
    Sense polarization
  Calculate the Direct e31 coefficient by dividing the surface strain into the sense polarization.
Dual e31 Measurement

## Slide 18
4x4 Mux
Tester
RETURN
DRIVE
SENSOR
Pogo 4
Pogo 3
Pogo 2
Pogo 1
Dual E31 Fixture
4 x 4 Mux
  The fixture has a built-in 4x4 multiplexer to allow Vision to connect its DRIVE, RETURN, SENSOR, or GND channels to any one of the four pogo pins contacting the cantilever capacitor pads during execution of a test.

## Slide 19
Measure Outer Capacitor
Tester
DRIVE
RETURN
  Drive outer capacitor
  Measure displacement
  Measure outer capacitor polarization
  Slow, 0.5 Hz (2s) or less

## Slide 20
Measure Inner Capacitor
Tester
DRIVE
RETURN
  Drive inner capacitor
  Measure displacement
  Measure inner capacitor polarization
  Slow, 0.5 Hz (2s) or less

## Slide 21
  The exact overlay of the shapes of the two loops indicates that there is a single negative coefficient that can multiply one to get the other:
Evaluate Results
Normalized Direct e31 Charge
Normalized Displacement Loop
Charge:
Displacement:
  Peak displacement: 5.5 µm.
  Peak Polarization: -0.002 µC/cm2
  Direct e31 Coefficient: -12.9 C/m2